## Slide 1: Therac-24
The Upshot

## Slide 2: Summary/Overview
Six patients received radiation overdoses during cancer treatment by a faulty medical linear accelerator, the Therac-25 unit
Overdoses caused by programming errors (that produced “race conditions”)
Case has led to advancements in systems safety (testing, computer control, reporting)
Industry response was inadequate
Prompting and input by regulatory agencies (FDA) and User Groups were instrumental in finding causes of and remediating the accidents

## Slide 3: Hager’s perspective
Hager knew that something was clearly wrong and attacked the problem by working with the operator after the April 11th incident to recreate the events leading to the overdose
The operator remembered the sequence of data entry procedures and repeated over and over the procedure until she could consistently reproduce the error
They correctly diagnosed the problem as a race condition produced by a flaw in the software programming.

## Slide 4: The Problem
Race Condition (From Huff and Brown)
  Produced the two deaths in ETCC
  Discovered and duplicated by Hager working with machine operator
Those with programming experience can consult Computing Cases for detailed explanation
Philosophers and forks
  Philosophers eat and think. They also share forks with those sitting on either side. They need two forks to eat properly. (They’re messy)
  To solve the problem of scarce resources, philosophers need to coordinate thinking and eating. Imagine Philosophers A, B, and C sitting at a round table and sharing the forks on either side of their plates.
    While Philosopher B is thinking, Philosophers A and C are eating
    While Philosopher B is eating (and using the forks on either side), Philosophers A and C are thinking.
  A race condition occurs when the coordination breaks down and Philosophers A, B, and C all simultaneously reach for their forks. If A and C get there first, B will starve. (Philosophers are notoriously unable to adapt to changing circumstances because they respond only to a priori conditions.)

## Slide 5: Disseminating the results
Having identified the problem, it is now possible to return to the chronology
Different sites began to share their experiences
  Operators held special sessions at conferences on their experiences
FDA required a CAP from AECL
  Corrective Action Plan

## Slide 6: FDA Pre-Market Approval
Class I
  “general controls provide reasonable reassurance of safety and effectiveness””
Class II
  “require performance standards in addition to general controls”
Class III
  Undergo premarket approval as well as comply with general controls
Used earlier Therac models to show “pre-market equivalence”
  But this covered over three key changes:
    removal of hardware safety controls,
    delegation of safety from hardware to software,
    No testing of additional programming for Therac-25 layered on programming for 6 and 20 units

## Slide 7: FDA couldn’t recall defective products
Ask for information from a manufacturer
Require a report from the manufacturer
Declare a product defective and require a corrective action plan (CAP)
Publicly recommend that routine use of the system on patients be discontinued
Publicly recommend a recall

## Slide 8: FDA’s powers were limited
Mostly persuasive
Operators and hospital physicists worked to fill in the gap by assembling dispersed information on the operating history of Therac-25 and putting together a coherent story to explain the patient complaints.

## Slide 9
Date | Event
May 1 1986 | Second ETCC victim dies
May 2 | FDA declares T-25 defective and their response (alterations) inadequate. FDA demands a CAP from AECL
June 13 | AECL produces first CAP
July 23 | FDA requests more information from AECL
August | First Therac user group meeting
August | First ETCC victim dies and family files lawsuit

## Slide 10
Date | Event
Sept 26, 1986 | AECL provides more information to FDA
Oct 30 | FDA requests still more information
Nov | Physicists and engineers from FDA’s CDRH conduct a technical assessment of T-25 unit in Canada
Nov 12 | AECL submits second CAP revision
Jan 17, 1987 | Second patient overdose at Yakima
Jan 19 | AECL issues hazard notification to users. Visually confirm position of turntable

## Slide 11
Date | Event
Jan 26, 1987 | Conference call between AECL quality assurance manager and FDA. AECL sends FDA revised test plan. AECL calls T users with instructions on how to avoid beam on when turntable is in field light position.
Feb 3 | AECL announces additional changes to T
February 1987 | Therac-25 units in Canada and US are taken out of service until AECL completes new CAP
March 5 | AECL completes third CAP. FDA (April) asks for more information)
April 13 | AECL sends update of CAP plus list of nine items requested by users at march meeting
May 1 | AECL sends fourth revision of CAP to FDA

## Slide 12
Date | Event
May 26 | FDA approves fourth CAP subject to testing and analysis
June 5 | AECL sends final test plans to FDA along with safety analysis
July | Third T-25 user group meeting
July 21 | AECL sends final (fifth) CAP revision to FDA
Jan 28, 1988 | Interim safety analysis report issued from AECL
Nov 3, 1988 | Final safety analysis report issued

## Slide 13: Therac-25 Concepts
Safety, Risk, and Informed Consent

## Slide 14: Safety
“A thing is safe if, were its risks fully known, those risks would be judged acceptable in light of settled value principles.” (Martin/Schinger, Engineering Ethics, 108)
Safety and risk are different sides of the same coin
  One is defined in terms of the other
“Settled value principles” makes safety a matter of public policy. Government plays a role. So does business. Most importantly, so do members of the public

## Slide 15: Public
“those persons whose lack of information, technical knowledge, or time for deliberation renders them more or less vulnerable to the powers an engineer wields on behalf of his client or employer”
  Michael Davis. Thinking Like An Engineer
The public is in an especially vulnerable position. They stand subject to the risk. But they do not participate in the project that generates the risk
The public has the right to free and informed consent.
  This right is vulnerable if risk information does not get to them, if the risk information is too complicated for them to appreciate, or no provisions have been taken to include them in the collective risk acceptability (=safety) decision.

## Slide 16: Risk
The other side of the coin
  Risk and safety are correlative and defined in terms of one another
“A risk is the potential that something unwanted and harmful may occur.” (MS 108)
Risk has four dimensions (assessment, management, perception, and communication)
Since risk is the probability of harm and probability implies uncertainty (lack of complete knowledge), the ethics of risk lies in how this uncertainty is communicated and distributed.
  For example, does a government regulatory agency approve a product unless it is proven harmful….
  Or does it withhold approval from a product until it is proven completely safe.
  In the first, the burden of uncertainty falls on the public exposed to risk, in the second on the manufacturer who can’t reap benefits from selling the uncertainly risky product.

## Slide 17: Risk Assessment
The scientific and exact process of determining the degree of risk
Animal Bioassays
  Animals exposed to risk fact at intense level for short period of time
  Projection from animal physiology to human physiology and from short term/intense exposure to long term/less intense exposure
Epidemological Studies
  Comparison between populations exposed to risk and populations not exposed to risk
  Search for significantly higher risk ratio. Three-to-one not generally significant. Six-to-one is significant
Ethics of Risk
  Since there is uncertainty in risk assessment, an ethical issue arises as to how that uncertainty is distributed

## Slide 18: Risk Communication
Results of risk assessment are technical and subject to different interpretations
Public has a right to informed consent vis a vis risk
  To consent to take a risk (or withhold consent) they must understand the risk and be able to make a coherent consent decision
This raises issues in risk communication
  Clear communication
  Comprehensive communication (not leaving out anything significant)
  Communication that takes into account the perspective from which the public will perceive the risk

## Slide 19: Risk Perception
The public perceives risk according to a clear perspective
This renders risk perception rational because predictable (to a certain extent)
Factors which influence public perception of a risk’s acceptability
  Voluntariness
  Expected benefits
  Control over risk
  Minimal dread factor
  Minimal unknown factor

## Slide 20: Risk Management
Political process of determining if a certain degree of risk is acceptable according to a community’s settled value principles
Value principles are identified via a process of deliberative democracy which respect the meta-norms of reciprocity, publicity, and accountability
Community’s identify small scale project for experimental analysis
  These validate settled values
  These also help to determine if larger scale action is acceptable

## Slide 21: Safety procedures to consider when developing systems dependent on software
Nancy Leveson
Safeware: System Safety and Computers
p. 552

## Slide 22: Lessons
Software specifications and documentation should not be an afterthought.
Rigorous software quality assurance practices and standards should be established.
Designs should be kept simple and dangerous coding practices avoided
Ways to detect errors and get information about them, such as software audit trails, should be designed into the software from the beginning

## Slide 23: Lessons
The software should be subjected to extensive testing and formal analysis at the module and software level; system testing alone is not adequate. Regression testing should be performed on all software changes.
Computer displays and the presentation of information to the operators, such as error messages, along with user manuals and other documentation need to be carefully designed.

## Slide 24: Lessons
Reusing software modules does not guarantee safety in the new system to which they are transferred and sometimes leads to awkward and dangerous designs.
Safety is a quality of the system in which the software is used; it is not a quality of the software itself.

## Slide 25: Sources
Nancy G. Leveson, Safeware: System Safety and Computers, New York: Addison-Wesley Publishing Company, 515-553
Nancy G. Leveson & Clark S. Turner, “An Investigation of the Therac-25 Accidents,” IEEE Computer, 26(7): 18-41, July 1993
www.computingcases.org (materials on case including interviews and supporting documents)
Sara Baase, A Gift of Fire: Social, Legal, and Ethical Issues in Computing, Upper Saddle River, NJ: Prentice-Hall, 125-129
Cranor, Carl. (1993). Regulating Toxic Substance: A Philosophy of Science and the Law. Oxford, UK: Oxford University Press
Mayo, D.G. and Hollander, R.D. (1991). Acceptable Evidence: Science and Values in Risk Management. Oxford, UK: Oxford University Press.
Chuck Huff and Richard Brown. “Integrating Ethics into a Computing Curriculum: A Case Study of the Therac-25”
  Available at www.computingcases.org (http://computingcases.org/case_materials/therac/supporting_docs/Huff.Brown.pdf) Accessed Nov 10, 2010